## Slide 1: Summary of 2nd Microwave Calibration Workshop
S. Schmuck, Workshop Participants

## Slide 2: Agenda Overview
Monday:
  ITER site visit
  Review of previous calibration report done by ITPA working group
  CTS: 2 talks (W7-X, ITER) & ways forward to be discussed
  Refl LFS : 1.5 talk (ITER, LHD)
Tuesday:
  ECE: 7 talks (SPARC, HL-3, WEST, ITER)
  ECE: ways forward to be discussed
  Joint data analysis scheme: 1 talk (ITER)
  Development of reference calibration discharge: 1 talk (ITER)
  Round table
  Summary and Closure

## Slide 3: Review of previous calibration report done by ITPA working Review of previous calibration report done by ITPA working group
Reports done in 2010 and 2011 to document diagnostic procedures which should be done in-lab, in commission phase, during and after plasma operation ITER_D_33ZRFR and I TER_D_3622CE:
  Past work must not be forgotten

## Slide 4: CTS
Talk for ITER:
  Explanation on 3.5 mm diameter of launching mirror mm to be clarified
    To be checked soon as CTS components are under manufacturing
  Measure of non-linearities: establish procedure
  Launcher line transmission throughput could be monitored
  Coupling of calibration signal launched into receiving line might differ from coupling during plasma operation

## Slide 5: Reflectometer
GA:
  Switches too slow with 100 ms to protect detector, fast protection possible by diode switches with 1 mus response time followed by slower shutter
  Protection implies:
    Down-time of probing plasma of at least 200 ms: real-time control affected but LFSR is only basic control diagnostic
    Alternative: switching between lines when one line is affected only? reducing down-time to 100 ms:
      Open question: All antennas may be affected at the same time but sensors would reveal if ECH would be present
  Perform frequency sweep to determine non-linearity:
    Calibration time per frequency could be improved: processing latency to be reduced by moving to perform analyses to FPGA
    Single calibration 3-5 s to few 100 ms with current synthesis could be improved via faster synthesiser
  Improve detection: ITER B is constant, first reflection confined in space in frequency
LHD:
  Background noise elimination: benefit for high-frequency reflectometer to remove plasma related noise floor

## Slide 6: ECE
LHD:
  Cross-calibration of heterodyne radiometer with TS
SPARC:
  Quasi-optical arrangement: alignment active maybe needed
  FTS should be there to make sure radiation is thermal

## Slide 7: ECE
HL-3:
  Compare results of Magnetic Difference Calibration Method with conventional
  absolute calibration pending on upgrade of diagnostic system
  Te estimated from TS taken to estimated ECE sensitivity: How can inconsistencies be spotted?
WEST:
  FTS system suggested
  Find explanation of 50% change of channel sensitivity and try to mitigate
  Calibration procedure to be checked for movement of equipment away from machine to calibration stand
  Help offered to WEST group for calibration approach

## Slide 8: ECE
ITER:
  Comparing results of different diagnostics: take into consideration differing temporal and spectral resolutions of TS, FTS and radiometer
  Size of hot source by place of front-end, not to calibrate diagnostic below 120 GHz?
    Antenna pattern to be checked hopefully via measurements at IO, IN-DA and US-DA sites
  Derived model and results FTS sensitivity can be shared
  Noise estimation from different sources: plasma, detector, post detection

## Slide 9: Reference Calibration Discharge
ITER joint data analysis scheme:
  Check objectively compliance of diagnostic assumption given sensitivity, line of sight coordinates
  Systematics can be found and quantified and hopefully be resolved
ITER reference calibration discharge:
  Reference pulses carried out JET and AUG until somepoint in time but schedule pressure removed these pulses which introduces uncertainties
  How to participate in the design of the reference pulse?
    Submit requirements to M. Kocan or other diagnostic TROs in the next months
  Impact statements to justify and defined calibration pulses as
    BEST also planning for reference pulses, collaboration was agreed and IO waiting for feedback
    s

## Slide 10: Suggestions
Workshops suggested:
  3rd MWCM in 2 years’
  Alpha particle dominated ITER plasma: how to deal with non-Maxwellian kinetic profiles: science department to be consulted
  Systematic differences on ECE/TS Te profiles: evidence, causes and counter measures in next year
  Reference pulse in next year

## Slide 11
THANKS FOR ALL YOUR CONTRIBUTIONS AND ATTENDANCE
SEE YOU SOON HERE AGAIN