## Slide 1: Calibrating the ECE diagnostic for the ITER Tokamak
Modeling the Overall Sensitivity of the Fourier Transform Spectrometers
Johannes van den Berg Stefan Schmuck
Notes: Project
Diagnostics

## Slide 2: Electron Cyclotron Emission (ECE)
From [1].
Emission spectrum
O mode
X mode
From [1].
Frequency
Pure O1
Pure X2
ITER ECE: Te from pure O1 and Pure X2
Notes: Rayleigh-Jeans approx.

## Slide 3
ITER’s ECE diagnostic
From [4].
Quasi-Optical Antenna
and Polarizer
2 Fourier Transform Spectrometers 2 Heterodyne Radiometers
4 Oversized Circular Waveguide Transmission Lines
Antenna
Transmission Line
FTS
B74

## Slide 4
Radial sightline
Oblique sightline
Sightlines and Polarizations
Mainly-X
Mainly-O
×4

## Slide 5
Feeding into diagnostic backends
2 Fourier Transform Spectrometers 2 Heterodyne Radiometers
70–1000 GHz
Spectral Domain
70–1000 GHz
220–340 GHz
X2
122–230 GHz
O1 or half-field X2

## Slide 6
Feeding into diagnostic backends
2 Fourier Transform Spectrometers 2 Heterodyne Radiometers
70–1000 GHz
Spectral Domain
70–1000 GHz
220–340 GHz
X2
122–230 GHz
O1 or half-field X2

## Slide 7
Antenna
Transmission Line
FTS
Plasma inference mode
From [1].
Emission spectrum
Absolute calibration
required!
– Even sensitivity
– Odd sensitivity
(accounts for imperfections)
Measured

## Slide 8
Antenna
Transmission Line
FTS
Calibration mode
Hot-Cold Calibration sources inside DSM
Known
Overall sensitivity
Noise!
Radiative temperature sensitivity
→

## Slide 10
10×–100× noise reduction
8 hours ~ 35 days

## Slide 11
Antenna
Transmission Line
FTS
Estimating Overall Sensitivity

## Slide 12
Antenna
Transmission Line
FTS
Estimating Overall Sensitivity
  Truncation
  Mirror Ohmic Loss
  Polarizer Loss
  Free space coupling
  Waveguide Ohmic Loss measurement
  MB Ohmic Loss
  MB Diffractive Loss
  Windows
  Gap
Already included in prototype sensitivity measurement
To be included in more complete analysis
  Sensitivity based on measurement with IN-DA FTS prototype

## Slide 13
Antenna
Transmission Line
FTS
Power divider
HR

## Slide 14
Antenna
Transmission Line
FTS

## Slide 15: Quasi-optical beam diffractivity
2D slice of a Gaussian beam propagating
Beam width calculation along the optical path using complex beam parameter → ray transfer matrices
A more thorough analysis would include beam astigmatism and higher order modes
Waveguide entrance
Source
Focusing mirror
Focusing mirror
Foc. Mir.

## Slide 16: Antenna components
Input: component positions and dimensions
We assume the (projected) components to be circular
  We conservatively take whichever diameter is smaller (significant for polarizer)
Per beam path:
2–3 mirrors in the US-DA front end (depending on operation modus)
2–3 mirrors in the IN-DA PSU (depending on polarisation)
PSU
DSM2 Equatorial Port 9
Projected
Mirrors & polarizer

## Slide 17: Aperture/truncation loss
Path | Power Loss [%] at |  |  | 
 | 100 GHz | 200 GHz | 500 GHz | 1000 GHz
Radial X-mode | 13.5 | 1.13 | 0.69 | 0.66
Radial O-mode | 13.5 | 1.14 | 0.70 | 0.66
Oblique X-mode | 12.6 | 1.11 | 0.69 | 0.66
Oblique O-mode | 12.6 | 1.12 | 0.70 | 0.66
Power coupled through an aperture into a GB of the same width:
Loss due to beam truncation and partial filling of antenna pattern by calibration source
Power missing from
antenna pattern:
No Aperture D (~25% Loss)

## Slide 18: Aside: Apertures are problematic
Misalignment critical for sensitivity / filling the antenna pattern below ~200 GHz
  >1% loss, indicating beams are starting to be truncated
Stray radiation due to multi-reflections inside closed spaces like DSM/PSU
Discrepancy between sensitivities for Calibration and Plasma operation due to shutter mirror and discrepancy between filling of antenna patterns (>2% difference below ~120GHz)

## Slide 19: Mirror Ohmic Loss
E- and H-plane losses: O-mode is mostly H-plane reflections, so has fewer Ohmic losses on the mirrors
The bonus resistivity due to surface rougness is taken as 1.2, a common value in literature. In reality, this factor is likely frequency dependant.
H-plane
E-plane

## Slide 20: Track polarization through Antenna and Transmission Line

## Slide 21: Grid Polarizer Power Transmission

## Slide 22: Antenna Overall

## Slide 23
Antenna
Transmission Line
FTS

## Slide 24: Transmission Line Ohmic Loss measurement
Short
Long
Path | Length [m] | Power Loss [%] at |  |  | 
 |  | 100 GHz | 200 GHz | 500 GHz | 1000 GHz
Radial X-mode | 41.4 | 75.6 | 81.7 | 92.3 | 98.2
Radial O-mode | 44.1 | 77.7 | 83.6 | 93.5 | 98.6
Oblique X-mode | 42.8 | 74.9 | 82.7 | 91.9 | 98.0
Oblique O-mode | 40.6 | 76.7 | 81.1 | 92.9 | 98.4

## Slide 25: Transmission Line Ohmic Loss measurement
Path | Length [m] | Power Loss [%] at |  |  | 
 |  | 100 GHz | 200 GHz | 500 GHz | 1000 GHz
Radial X-mode | 41.4 | 75.6 | 81.7 | 92.3 | 98.2
Radial O-mode | 44.1 | 77.7 | 83.6 | 93.5 | 98.6
Oblique X-mode | 42.8 | 74.9 | 82.7 | 91.9 | 98.0
Oblique O-mode | 40.6 | 76.7 | 81.1 | 92.9 | 98.4
Short
Long
Surprisingly high! Overmoded circular waveguides are expected to have low losses (<1 dB/km) [Thumm 2002]

## Slide 26: Mitre Bends
TE₀₁ eigennumber from Patel 2010
Diffractive Losses:
Ohmic Losses are same as Mirrors!
H-plane
E-plane

## Slide 27: Losses Overall
High losses from TL realistic?
Transmission
Line
Transmission
Line
Path | Power Loss [%] at |  |  | 
 | 100 GHz | 200 GHz | 500 GHz | 1000 GHz
Radial X-mode | 80.3 | 82.7 | 92.8 | 98.3
Radial O-mode | 82.2 | 84.6 | 93.9 | 98.7
Oblique X-mode | 81.1 | 83.7 | 93.4 | 98.6
Oblique O-mode | 79.6 | 82.2 | 92.4 | 98.2

## Slide 28
Antenna
Transmission Line
FTS

## Slide 29: Backend sensitivity
Uncertainty covariance
Obtained via Bayesian approach
(see talk by S. Schmuck)

## Slide 30: Overall sensitivity
Very high uncertainty
at low frequencies
But we have an order of magnitude estimate
Expected difference
interferograms:
High variance past 15mm needs investigation

## Slide 31: Measurement Uncertainties of Sensitivity
Noise level of IN-DA FTS for individual interferogram at = 9.8 mV
When for each calibration source interferograms are acquired, measurement noise of difference interferogram reads
For FTS shaker frequency = 25 Hz and total available integration time :
  Sensitivities of forward & backward FTS mirror motions different: =
  Sensitivities of forward & backward FTS mirror motions same: =
When sensitivity estimated conventionally from data vector with data points via
then Gaussian noise in data domain translates into Gaussian noise in spectral domain
To reduce noise in spectral domain:
  Linear reduction: increase of probed sources
  Linear reduction: reduce measurement noise
  Square-root reduction: take more data per interferogram with with
  Square-root reduction: increase integration time
  Square-root of 2: symmetric power coupling wrt mirror movement yielding 1 if not

## Slide 32: Integration Time to Reach Relative Uncertainty < 10%
5.3 T: X2
Assumptions to estimate relative uncertainty wrt integration time :
  No windows, expansion joints, nor aperture D
  Antenna pattern filled sufficiently by calibration sources
  Emission spectra of calibration sources perfectly known
  ( C, C)
Relative uncertainty follows
  Problematic domain below 180 GHz: 24 – 48 h integration time needed
Te profile domain affected most:
  For full field of B = 5.3 T, LFS determined from O1 less certain
  For half field of B = 2.65 T, LFS determined from X2 less certain
Sensitivity differing
for-& backward
motions FB=1
Sensitivity same
for-& backward
motions FB=1/2
5.3 T: O1
Notes: Important ranges of frequency band

## Slide 33: Experimental/Design Suggestions
Measurement of Transmission Line at multiple lengths to pin down contributions of higher order modes, misalignments, etc.
  Attenuation much larger than literature needs to be explained.
Investigate impact of antenna component misalignment on sensitivity (below 200 GHz)
Characterize impact of shutter mirror and incomplete antenna pattern on validity of calibration (below 120 GHz) vs

## Slide 34: Modeling Improvements
Add windows, gap losses, and power splitter to sensitivity estimate. (straightforward)
Refine Bayesian method. Uncertainty in O1/X2 domain is very high.
  >15mm tails on interferogram seem spurious.
Analyse both antenna and waveguide while taking into account higher order modes. The waveguide is overmoded!

## Slide 35: Sources
[1] Schmuck, Stefan. Bayesian Inference for Microwave Diagnostics at Joint European Torus. Diss. KTH Royal Institute of Technology, 2022.
[2] Schmuck, Stefan, et al. "Electron cyclotron emission measurements on JET: Michelson interferometer, new absolute calibration, and determination of electron temperature." Review of Scientific Instruments 83.12 (2012).
[3] Schmuck, Stefan, and Jakob Svensson. "Fourier spectroscopy: a Bayesian way." International Journal of Spectroscopy 2017.1 (2017): 9265084.
[4] Pastor, Orlando. “Introduction to 55.F1” ITER IDM: 7W4TZC.

## Slide 36
Path | Length [m] | Attenuation at |  | 
 |  | 100 GHz | 500 GHz | 1000 GHz
Radial X-mode |  | 16% | 1.2% | 0.66%
Radial O-mode |  |  |  | 
Oblique X-mode |  |  |  | 
Oblique O-mode |  |  |  | 

## Slide 39
Antenna
Transmission Line
FTS